CHAPTER – IMPLEMENTATION CHALLENGES AND SOLUTIONS

Problems, Errors and Challenges Faced During Project Development

During the development of the Smart Diabetic Foot Ulcer Monitoring System, several practical challenges were encountered in hardware integration, sensor calibration, communication, and data processing. The following are the major issues faced and the solutions implemented.

1. FSR Sensor Not Giving Output

Problem:

Initially, the FSR sensors were not producing any readings (always showing 0 in Serial Monitor).

Cause:

- Improper circuit connection

- Incorrect resistor value in voltage divider

- Weak pressure sensitivity

Solution:

- Implemented proper voltage divider circuit

- Replaced 10kΩ resistor with 4.7kΩ resistor

- Verified connections and grounding

2. Inconsistent Pressure Values

Problem:

The sensor values were not matching actual pressure applied.

Cause:

- FSR sensors are non-linear

- Different sensors give different outputs for same force

Solution:

- Avoided absolute calibration

- Used relative pressure analysis

- Collected dataset from multiple users instead of fixed thresholds

3. Serial Monitor Output Too Fast

Problem:

Values were changing too fast and unreadable.

Solution:

- Added delay:

delay(1000);

- Reduced baud rate noise issues

4. Moisture Sensor Not Working

Problem:

Soil moisture sensor did not give proper readings for skin.

Cause:

- Sensor designed for soil, not human skin

Solution:

- Removed hardware moisture sensor

- Introduced Skin Moisture Index (SMI) using:

  - Skin temperature

  - Environmental temperature

  - Humidity (API)

5. OpenWeather API Not Working in ESP32

Problem:

Environmental data (temperature & humidity) always returned 0.

Cause:

- HTTPS request issues in ESP32

- Memory limitations

Solution:

- Shifted API handling to Python server

- ESP32 used only for sensor data transmission

6. WiFi IP Address Changing

Problem:

ESP32 IP address changed when reconnecting to WiFi.

Solution:

- Implemented Static IP configuration

IPAddress local_IP(192,168,31,200);

7. ESP32 Not Showing IP Address

Problem:

IP address not printed in Serial Monitor.

Cause:

- WiFi not connected properly

- Missing debug prints

Solution:

- Added proper connection loop:

while (WiFi.status() != WL_CONNECTED)

- Printed using:

Serial.println(WiFi.localIP());

8. Connection Loss Between ESP32 and Python

Problem:

Python server failed to receive data when connection dropped.

Solution:

- Added timeout handling

- Displayed warnings:

  - "LEFT SHOE DISCONNECTED"

  - "RIGHT SHOE DISCONNECTED"

9. Google Colab Not Working for Data Collection

Problem:

Socket connection failed in Colab.

Cause:

- Colab cannot access local network devices

Solution:

- Ran Python server in local system instead of Colab

10. Data Loss During Collection

Problem:

If program stopped, collected data was lost.

Solution:

- Implemented real-time saving to CSV/Excel

- Data saved continuously after each record

11. Handling Zero Values in Dataset

Problem:

Some sensors showed 0 values during walking.

Clarification:

- This is valid behavior

- Foot does not press all sensors at same time

Solution:

- Removed rows only when all sensors = 0

- Kept partial zero values for real analysis

12. Difficulty in Model Training

Problem:

Hard to define threshold for ulcer detection.

Solution:

- Changed approach:

  - Instead of classification → used health percentage

- Used:

  - Pressure distribution ratios

  - SMI

  - Body parameters

13. UI Not Showing Proper Status

Problem:

UI showed “Disconnected” even when devices were connected.

Cause:

- Wrong IP / network mismatch

Solution:

- Ensured:

  - Same WiFi network

  - Correct static IP

  - Proper socket connection

14. Noise and Random Values in Sensors

Problem:

Sensor values fluctuated randomly.

Solution:

- Used:

  - Averaging technique

  - Stable delay intervals

- Ignored minor fluctuations

15. Synchronization Between Two Shoes

Problem:

Left and right data were not synchronized.

Solution:

- Read both data within same loop

- Processed only when both available

Conclusion of Challenges

Despite multiple hardware and software challenges, the system was successfully developed by:

- Using adaptive approaches instead of rigid calibration

- Shifting complex tasks (API, ML) to Python

- Implementing reliable communication between ESP32 and server

- Designing a robust data collection and processing pipeline

These challenges and solutions provide valuable insights for future improvements and real-world deployment of smart healthcare systems.